## Slide 1: Guide to Tuolumne River Stream Data Yosemite, California
Jessica Lundquist
Gwyn Perry, Eric Keenan
Civil and Environmental Engineering, University of Washington
Harrison Forrester, Jim Roche
Yosemite National Park
Water Years 2002 to 2015
Published 2016

## Slide 2: Table of Contents
Overview map 3
Instruments and Installations 4
Processing Steps for Stage 6
Data Types 9
Rating Curve Process 10
Photos, Cross-sections, Rating curves by site (UW sites)
  Lyell Fork Below Maclure 13
  Lyell Fork Above Twin Bridges 18
  Lyell Fork Below Twin Bridges 24
  Dana Fork at the Lodge 28
  Dana Fork at Bug Camp 32
  Tuolumne at Highway 120 37
  Delaney Creek 44
  Budd Creek 48
Info on Other Stream Sites
  Grand Canyon of the Tuolumne NA
  Falls Creek NA
  Tuolumne at Hetch Hetchy NA

## Slide 3: Overview of Yosemite Hydroclimate Dataset
Distributed temperature data from Lundquist and Cayan 2007
Nested streamflow data for Tuolumne contributing to Hetch Hetchy (the ASO LiDAR domain)
Full met data to force a hydro model

## Slide 4
Brief History of “Wilderness Stream Gauging”
Aug 2001: Solinst and Seabird in anchors
2015: Most sites upgraded to vented pressure transducer
Aug 2005: Switch to wilderness stilling tube
Staff plate
Stilling tube
Instruments and Installations

## Slide 5
Installation Type | Anchored Solinst | Solinst in Stilling Tube | Vented Pressure Transducer
Description | Instrument in a PVC pipe inside a concrete anchor, which is cabled to a tree, bridge, or culvert. | Instrument in PVC pipe inserted in vertical pipe attached to the streambed and bank with rebar; with cord for downloading instrument. | Same as stilling tube but with data cord connected to a data logger box (typically hidden in a tree) and another cord open to the atmosphere.
Instrument Used | Solinst Levelogger | Solinst Levelogger | Druck1 Or Campbell Scientific CS450 PT2
Instrument Specs/Accuracy | Levelogger Model 3001: 0.1C temp accuracy, ±0.5 cm pressure/depth accuracy; temperature compensated over the range of -10 to 40C; drift of 0.1% of the full range (±0.5 cm for a 5 m model, used here). | Levelogger Edge and Gold: Temp accuracy ± 0.05C Pressure ± 0.05% of FS (for 5 m model, this would be ±0.25 cm); Manufacturer states clock accurate to 1 minute per year, but 20 minutes of drift per year was typically observed | Druck: 0-5 PSI Range, 0.25% accuracy; CS450: 0-7.25 PSI Range, 0.1% accuracy;
Processing steps required | 1) subtract off atmospheric pressure; 2) correct for offsets in instrument location; 3) check for instrument drift; 4) develop rating curve | 1, 3, and 4 | 3 and 4
Total error estimates in stage (Note that these are worst case scenarios) | Up to ± 3 to 4 cm, with ± 2 cm due to summed instrument accuracy and drift for both stream and barometric instruments; and ± 1 to 2 cm more due to uncertainty in instrument location | Up to ± 2 cm due to summed instrument accuracy and potential drift for both stream and barometric instruments | Up to ± 0.5 cm due to summed instrument accuracy and potential drift
Error in estimated discharge* | ± 0.92 m3 s-1 to ± 1.24 m3 s-1 (14-19%) | ± 0.61 m3 s-1 (9%) | ± 0.15 m3 s-1 (2%)
*Using Lyell Fork Twin Bridges summer flow, 0.7 m stage, as an example

## Slide 6
Atmospheric pressure (the green line) is subtracted from the raw instrument pressure (the red line) to obtain pressure due to water overlying the instrument (the blue line).
Processing Steps for Stage
1. Subtract off atmospheric pressure (example from Dana Fork Bug Camp)

## Slide 7
Raw pressure from multiple years (each color represents a different instrument) illustrates how offsets are often required when instruments are replaced.
Manual stage measurements (black circles) provide guidance for offsets to adjust the barometrically-corrected water pressure values.
Stage = barocorrected pressure + offset
Raw pressure data
Processing Steps for Stage
2 & 3. Correct for offsets in instrument location and check for instrument drift (example from Delaney Creek)

## Slide 8
Measurements used to create rating curve (note 3 outliers marked here were not used)
Discharge timeseries: Blue is meadow record provided here; Red is independently generated record from slightly upstream.
Meadow (blue) is prone to ice jams.
Processing Steps for Stage
4. Develop rating curve (example from Delaney Creek)

## Slide 9: Data Types:
Raw data: pressure recording, including both weight of water in stream and atmosphere; pieced together from multiple instruments to create continuous record in the water
Baro-corrected data: Atmospheric pressure subtracted off the raw timeseries
Offset: Added to baro-corrected data to eliminate times when instrument moved and when different instruments had different local biases: corrected to match stage datum when and where available
Stage from instrument (in manual measurement files): This is baro-corrected data + offset. Used to create rating curve.

## Slide 10: Rating Curve Development
Pre-process manual data (as shown above)
  Include 10% uncertainty in discharge
Develop prior rating curve parameters from survey data, with substantial uncertainty values, using BaRatin (Le Coz et al. 2014), to fit the form Q=a(h - b)c
  Section Control: low flow acts like a rectangular weir
    Inputs for determining prior “a” =
      Discharge Coefficient (Cr): default = 0.4 ± 0.05
      Width of “weir”, perpendicular to flow direction (Br)
    Inputs for determining prior “b”
      Average elevation of “weir” crest
    Inputs for determining prior “c”
      Default: 1.5 ± 0.05
  Channel Control: mid flow acts like a rectangular channel, use the Manning–Strickler equation
    Inputs for determining prior “a” =
      Slope of channel bed (Se)
      Channel width (B)
      Roughness coefficient (Ks): 20 ± 5 for all channels to be conservative
    Inputs for determining prior “b”
      Average elevation of Channel bottom
    Inputs for determining prior “c”
      Default: 1.67 ± 0.05
  Channel + Floodplain Control: high flow acts like the sum of the channel and floodplain (sum of 2 channel controls)

## Slide 11: Rating Curve Development (continued)
Use the BaRatin MCMC routine (see LeCoz et al. 2014 and references therein) to use manual stage and discharge observations, along with their estimated uncertainty (10%), to determine the posterior rating curve parameters and control segment breaks to fit the following piecewise power function:
  Where Nrange is the number of stage ranges
  Ncontrol is the number of hydraulic controls
  Kr is the upper water level at stage range r
  M is the hydraulic control matric (M(r,j) = 1 if hydraulic control j is active in stage range r
  h is stage
  a, b, and c are the fitted parameters for stage range j
Outputs include the rating curve equation for each stage range, as well as a look up table of values for the upper and lower limits of 95% confidence intervals (which vary by stage and are determined as part of the MCMC process)
These are illustrated for each site in the pages that follow.
For comparison, a single equation of the form Q=a(h - b)c was also calculated for each site using transformed least squares (Rantz et al. 1982), and this equation and its 95% confidence intervals are also shown

## Slide 12: Specific Sites:
Lyell Fork below Maclure
Dana Fork of the Tuolumne River
Lyell Fork of the Tuolumne River
Tuolumne River at Highway 120
Delaney Creek

## Slide 13
Lyell Fork below Maclure
40 m
JMT/PCT Bridge
site
(stilling tube)
control
ice jams
Overall, rating curve is good, but there is greater uncertainty in the datum before 2008, and there are occasional ice jams.

## Slide 14: Lyell Fork below Maclure
Latitude: 37.777
Longitude: -119.261
Datum: NAD83
Elevation: 2940 m
Instrument location
View upstream from instrument
Flow

## Slide 15
Based on photographs estimate two cross-sections:
  Channel Control
    Width: 6 m +- 2 m
    Slope: 0.04 +- 0.02
  Section Control
    Width: 3 m +- 3 m (unknown cross-section)
    Height of control break: 3 m +- 3 m (unknown cross-section, maximum depth measured is 6 m)
Lyell Fork below Maclure Creek

## Slide 16
Lyell Fork below Maclure Creek
Least Squares (stage in ft, Q in cfs):
Q = 14.9123*(H – 6)^2.5323
BaRatin (stage in m, Q in cms):
2.25m < H <= 6.044m
Q = 11.609*(H –2.0443)^1.6415
H <2.25
Q =4.9324*(H – 1.9357)^1.4917
Posterior Height of Section Break between Section Control and Channel Control: 2.25 m
Posterior Height of Section Break between Section Control and Channel Control: 2.25 m

## Slide 17
Lyell Fork below Maclure Creek
Log-Log View
Posterior Height of Section Break between Section Control and Channel Control: 2.25 m
Posterior Height of Section Break between Section Control and Channel Control: 2.25 m

## Slide 18: Lyell Fork at Twin Bridges
100 m
Latitude: 37.869
Longitude: -119.331
Datum: NAD83
Elevation: 2640 m
Twin Bridges
Upper site
(stilling tube)
Lower site
(anchor)
Bedrock control
The downstream site was installed first (WY 2002), with the upstream site designed to replace it starting WY 2006. Calculated discharge agrees very well during the 2006-08 period of overlap, with the exception of winter and early spring ice jams. For 2009 on, the downstream site was less frequently maintained, and the upstream site record should be used.
Downstream site was subject to ice jams

## Slide 19
Upstream instrument
Low flow control
Discharge measurement at bank full flow (18 June 2003)
Flow
Flow
Bedrock control
Lyell Fork abv Twin Bridges
Latitude: 37.869
Longitude: -119.331
Datum: NAD83
Elevation: 2640 m

## Slide 20
Note: On 17 July 2015 at 9:15 am local time, the depth of water over the bed-rock control was 0.68 ft and over the gravel-bar control was 1.50 ft. Gauge height measured as 10 ft – distance from top of stilling tube to water (0.78 ft) at that time was 9.22 ft; making gauge height 8.54 ft when bedrock control matters and 7.72 ft when gravel-bar control matters – need to check new references to my stage measurements. Stage reference has been hard here.
Lyell Fork above Twin Bridges

## Slide 21
Chart: Total Width (m)
| 1.9812 | 3.048 | 3.9624 | 4.8768 | 5.7912 | 6.7056 | 7.619999999999996 | 8.5344 | 9.448799999999998 | 10.3632 | 11.2776 | 12.192 | 13.1064 | 14.0208 | 14.9352 | 15.8496 | 16.764 | 17.6784 | 18.5928 | 19.5072 | 20.4216 | 20.4216 | 21.33600000000001 | 21.33600000000001 | 22.2504 | 22.2504 | 23.1648 | 23.1648 | 24.0792 | 24.9936 | 25.908 | 26.8224 | 27.73680000000001 | 28.65120000000001
 | 0.8096 | 0.56576 | 0.47432 | 0.38288 | 0.38288 | 0.38288 | 0.38288 | 0.41336 | 0.47432 | 0.5048 | 0.5048 | 0.56576 | 0.59624 | 0.62672 | 0.53528 | 0.53528 | 0.5048 | 0.47432 | 0.44384 | 0.38288 | 0.23048 | 0.23048 | 0.2 | 0.2 | 0.23048 | 0.23048 | 0.23048 | 0.23048 | 0.3524 | 0.3524 | 0.62672 | 0.6572 | 0.6572 | 0.59624
Instrument
Channel
Floodplain
Flow
Slope: 0.0003 m/m
River Right
Instrument
Lyell Fork above Twin Bridges
Cross-Section at Instrument Site with Stage Measurement Datum
Height of Section Break between Channel Control and Floodplain/Channel Control: 0.8 m +- 0.2m
Height of Section Break between Section Control and Channel Control: 0.6 m +-0.2m
Width used for Channel Control (used same width for floodplain): 25 m +- 5m
Width used for Section Control: 15 m +- 5m
Note significantly different scales for x and y axes.

## Slide 22
Posterior Height of Section Break between Section Control and Channel Control: 0.54 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 0.80m
Posterior Height of Section Break between Section Control and Channel Control: 0.54 m
Lyell Fork above Twin Bridges
BaRatin (stage in m, Q in cms):
0.79875m < H
Q =30.94*(H –0.3023)^1.6658
0.540748m < H < 0.79875m
Q =35.6095*(H – 0.343)^1.662
H < 0.540748m
Q =11.5733*(H –0.1873)^1.512
Least Squares (stage in cm, Q in cfs):
Q = 0.00965*(H – 10)^2.4474
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 0.8m

## Slide 23
Posterior Height of Section Break between Section Control and Channel Control: 0.54 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 0.8 m
Lyell Fork above Twin Bridges
Posterior Height of Section Break between Section Control and Channel Control: 0.54 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 0.8 m
Log-Log View

## Slide 24
Lyell Fork blw Twin Bridges
Latitude: 37.869
Longitude: -119.331
Datum: NAD83
Elevation: 2640 m
Instrument

## Slide 25
Chart: Station (m)
| 0.0 | 0.4572 | 2.80416 | 3.2004 | 3.2766 | 7.07136 | 9.81456 | 12.10056 | 13.4112 | 14.99616 | 15.60576 | 15.66672 | 17.73936 | 18.34896000000001 | 18.40992 | 19.56816 | 20.78736 | 21.09216
 | 2.6853 | 2.094 | 1.7343 | 1.5514 | 1.3076 | 0.57912 | 0.44806 | 0.37186 | 0.19202 | 0.11887 | 0 | 0.9144 | 2.158 | 1.9111 | 2.5512 | 1.1887 | 1.9934 | 2.8804
Height of Section Break between Channel Control and Floodplain/Channel Control: 2.55 m +- 0.2m
Height of Section Break between Section Control and Channel Control: 0.91 m +- 0.2m
Width used for Channel Control: 21 m +- 5
Width Used for Floodplain/Channel Control: 25 m +- 5m
Width used for Section Control: 10 m +- 1m
Slope: 0.0111 m/m
River Right
Instrument
Lyell Fork below Twin Bridges
Note significantly different scales for x and y axes.

## Slide 26
Posterior Height of Section Break between Section Control and Channel Control: 1.08 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 2.4 m
Posterior Height of Section Break between Section Control and Channel Control: 1.08 m
Height above which rating curve does not apply (ice jam): 2.5m
Lyell Fork below Twin Bridges
BaRatin (stage in m, Q in cms):
2.425m < H <= 2.5m
Q = 103.003*(H –1.6484)^1.671
1.08171m < H < 2.425m
Q = 25.193*(H –0.6297)^1.6848
H < 1.08171m
Q =9.3003*(H – 0.28214)^1.526
Least Squares (stage in cm, Q in cfs):
Q = 0.0043*(H – 10)^2.4177

## Slide 27
Posterior Height of Section Break between Section Control and Channel Control: 01.08 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 2.4 m
Height above which rating curve does not apply (ice jam): 2.5m
Lyell Fork below Twin Bridges
Log-Log View

## Slide 28: Dana Fork, Tuolumne Lodge and Bug Camp
Lodge site (anchor)
Bug Camp site (stilling tube)
bridge
Latitude: 37.876 Longitude: -119.333
Datum: NAD83
Elevation: 2800 m
Latitude: 37.877
Longitude: -119.338
Datum: NAD83
Elevation: 2640 m

## Slide 29
Chart: Station (m)
| 0.67056 | 1.6764 | 4.267199999999995 | 6.7056 | 10.0584 | 10.3632 | 10.91184 | 11.7348 | 11.7348 | 11.97864 | 12.71016 | 12.8016 | 14.0208 | 19.2024 | 21.9456
 | 1.8349 | 1.0485 | 0.87173 | 0.63094 | 0.3109 | 0.042672 | 0.033528 | 0 | 0.17983 | 0.10668 | 0.49682 | 0.94183 | 1.3777 | 1.4996 | 1.5301
Maximum height of measured stage: 1.4m
Height of Section Break between Section Control and Channel Control: 0.5 +- 0.2m
Width used for Channel Control: 12.5m +- 5m
Width used for Section Control: 3m +- 1m
River Right
Instrument
Dana Fork near Tuolumne Lodge
Note significantly different scales for x and y axes.

## Slide 30
Dana Fork near Tuolumne Lodge
BaRatin (stage in m, Q in cms):
0.456733m =< H
Q = 31.249*(H –0.4419)^1.6651
H < 0.456733m
Q =4.0361*(H –0.421)^1.4915
Least Squares (stage in ft, Q in cfs):
Q = (103.705*(H – 36.704)*0.03281)^2.355
Posterior Height of Section Break between Section Control and Channel Control: 0.46 m

## Slide 31
Dana Fork near Tuolumne Lodge
Posterior Height of Section Break between Section Control and Channel Control: 0.46 m
Log-Log View

## Slide 32
Staff plate installed 2012
Photo: 7/16/2015
corded Solinst (Silver model) inside stilling tube, installed 2006
Flow
Flow
Dana Fork at Bug Camp

## Slide 33
Chart: Station (m)
| 0.0 | 1.299810755437293 | 2.594188312080882 | 3.586267697300172 | 4.28145886356117 | 4.816802673362648 | 5.027701761011027 | 6.424255442725681 | 7.795823560931207 | 8.25493621957205 | 10.60295076855807 | 11.35009400848862 | 11.40024267258994 | 12.11712573141925 | 12.21853182624422 | 13.06501105955797 | 13.10014202175832 | 13.99135736748855
Elevation (m) | 1.021 | 0.294 | 0.174 | 0.084 | 0 | 0.092 | 0.017 | 0.094 | 0.298 | 0.26 | 0.309 | 0.364 | 0.323 | 0.286 | 0.336 | 0.393 | 0.426 | 0.887
Height of Section Break between Channel Control and Floodplain/Channel Control: 0.9 m +- 0.2m
Height of Section Break between Section Control and Channel Control: 0.32 m +- 0.2m
Width used for Channel Control: 12.7 m +- 5m
Width Used for Floodplain/Channel Control: 15 m +- 5m
Width used for Section Control: 6.5 +- 1m
Slope: 0.033 m/m
River Right
Instrument
Channel
Floodplain
Flow
Dana Fork at Bug Camp
Note significantly different scales for x and y axes.

## Slide 34
Posterior Height of Section Break between Section Control and Channel Control: 0.36 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 0.98 m
Posterior Height of Section Break between Section Control and Channel Control: 0.36 m
Dana Fork at Bug Camp
BaRatin (stage in m, Q in cms):
0.976552m < H
Q = 41.69*(H – 0.79033)^1.692
0.357608m < H < 0.976552m
Q = 37.16*(H – 0.2854)^1.6395
H < 0.357608m
Q =8.822*(H – 0.2066)^1.5183
Least Squares (stage in ft, Q in cfs):
Q = 68.09*(H –0. 5)^2.5

## Slide 35
Posterior Height of Section Break between Section Control and Channel Control: 0.36 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 0.98 m
Dana Fork at Bug Camp
Log-Log View

## Slide 36
Dana near Bug Camp, 26 July 2003
Flow
Instrument location
Dana Fork at Bug Camp
Substantial uncertainty exists in the rating curve at the lodge location (see confidence intervals). Recommended to use the Bug Camp location data when possible.

## Slide 37: Tuolumne at Highway 120
Dana Fork
Downstream section control
Lyell Fork
Staff gauge and instrument location
Highway 120
Latitude: 37.876
Longitude: -119.355
Datum: NAD83
Elevation: 2600 m

## Slide 38
Flow
Original install in 2001, anchored solinst and/or Seabird at the base of this abutment.
Flow
Staff gauge, installed in 2006
Corded pressure transducer, installed in 2006
Tuolumne River at Highway 120

## Slide 39
Flow
Flow
Upstream of the bridge, photo taken 23 September 2003
Downstream of the bridge
Tuolumne River at Highway 120

## Slide 40
Chart: Station (m)
| 0.6096 | 4.267199999999995 | 6.096 | 10.0584 | 13.4112 | 16.15440000000001 | 17.0688 | 18.288 | 20.7264 | 21.9456 | 22.5552 | 22.86 | 23.7744 | 26.0604 | 26.51760000000001 | 26.51760000000001 | 27.432 | 27.8892 | 29.8704 | 32.004 | 32.91840000000001 | 33.528 | 34.13760000000001 | 35.052 | 37.1856 | 39.9288 | 45.11040000000001
Tuolumne Bridge x-section | 3.8191 | 2.2159 | 2.2068 | 1.4874 | 1.7435 | 1.6398 | 1.332 | 1.0424 | 0.762 | 0.49682 | 1.0455 | 0.42672 | 0.22555 | 0.26213 | 0.31394 | 0 | 0.042672 | 0.78638 | 0.62179 | 0.35966 | 0.41453 | 0.070104 | 0.42367 | 0.085344 | 0.94183 | 2.4872 | 3.173
Width Used for Floodplain/Channel Control: 45 m +- 5m
Width used for Channel Control: 36 m +- 5m
Height of Section Break between Channel Control and Floodplain/Channel Control: 2.48 m +- 0.2m
Height of Section Break between Section Control and Channel Control: 1 m +- 0.2m
Width used for Section Control: 22 m +- 2m
Slope: 0.0405 m/m
Instrument #1
Instrument #2
River Right
Tuolumne River at Highway 120
Note significantly different scales for x and y axes.

## Slide 41
Posterior Height of Section Break between Section Control and Channel Control: 0.76 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 2.5 m
Height above which rating curve does not apply (ice jam): 2.5 m
Posterior Height of Section Break between Section Control and Channel Control: 0.76 m
Tuolumne River at Highway 120
BaRatin (stage in m, Q in cms):
0.764609m < H <= 2.499m
Q = 18.007*(H – 0.7169)^1.713
H < 0.764609m
Q = 34.099*(H – 0.7460)^1.468
Least Squares (stage in cm, Q in cfs):
Q = 0.188488*(H – 70)^1.75063

## Slide 42
Posterior Height of Section Break between Section Control and Channel Control: 0.76 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 2.5 m
Height above which rating curve does not apply (ice jam): 2.5 m
Tuolumne River at Highway 120
Log-Log View

## Slide 43
Photo on 29 July 2015; red shows gravel bar section control. 16:00 on 29 July 2015: Gage height of zero flow 2.79 ft – 0.86 ft = 1.93 ft
Tuolumne River at Highway 120
The rating curve at this location has greater confidence than at many of the other sites.

## Slide 44: Delaney Creek
Meadow site (stilling tube)
Trail site
(stilling tube)
Tuolumne River
Trail to Glen Aulin
Latitude: 37.883
Longitude: -119.381
Datum: NAD83
Elevation: 2600 m
Meadow site only is available at this time. Notice frequency of ice jams.

## Slide 45
Chart: Station (m)
Maximum Stage Depth for which Rating Curve and Confidence Bounds are valid: 1.2 m
Floodplain
Channel
Instrument
Width used for Section Control: 6 m +- 1m
Flow
Slope: 0.00422 m/m
River Right
Instrument
Delaney Creek
Note significantly different scales for x and y axes.

## Slide 46
Posterior Height of Section Break between Section Control and Channel Control: 0.37 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 1.07 m
Height above which rating curve does not apply (ice jam): 1.2 m
Posterior Height of Section Break between Section Control and Channel Control: 0.37 m
Delaney Creek
BaRatin (stage in m, Q in cms):
1.07m < H <= 1.2m
Q = 41.69*(H – 0.79033)^1.692
0.369561m < H < 1.07m
Q = 6.157*(H – 0.19283)^1.685
H < 0.369561m
Q = 2.903*(H – 0.12934)^1.525
Least Squares (stage in cm, Q in cfs):
Q = 0.0037*(H – 6)^2.3618

## Slide 47
Posterior Height of Section Break between Section Control and Channel Control: 0.37 m
Posterior Height of Section Break between Channel Control and Floodplain/Channel Control: 1.07 m
Height above which rating curve does not apply (ice jam): 1.2 m
Delaney Creek
Log-Log View

## Slide 48: Budd Creek
Upper site
(stilling tube)
Lower site
(anchor)
Culvert under Highway 120
Latitude: 37.873
Longitude: -119.382
Datum: NAD83
Elevation: 2600 m

## Slide 49
culvert
Highway 120
Upstream instrument (stilling tube)
Staff plate
stream
Downstream instrument (concrete anchor)
Budd Creek
Budd Creek has a long record and two instruments, but due to a shifting channel bed, backwater effects from the culvert, and shifts in instrument location, the rating curve has high uncertainty. We used least squares (as in Rantz et al. 1982) and do not provide 95% confidence values for this site.

## Slide 50
Budd Creek
The upstream site serves as the principal record, except in the case of known ice jams or backwater affects. Occasionally water levels rise above the top of the culvert, therefore initiating higher recorded stage values which may not agree with rating curve estimated discharge values. In the event of backwater or ice jam events the downstream record was used.
The upstream site has a corded Solinst pressure transducer located inside of a stilling tube. Just downstream of the stilling tube is a white staff plate designed to standardize periodic manual stage measurements. These manual stage measurements were used to correct instrument drift in the Solinst pressure transducer.
Backwater from the culvert and changing downstream sediment make this site not work well with the BaRatin method.
Upstream of culvert
Staff plate, used as stage datum
Stilling tube
Flow

## Slide 51
Budd Creek
Downstream of culvert
The downstream Solinst pressure transducer is located in concrete anchor. The anchor occasionally is moved during high flows, thus requiring an estimated offset in order to provide a consistent record. The upstream of culvert site therefore provides a preferred record for Budd Creek.
Flow
Flow
Flow
Flow

## Slide 52
Rating curve for Budd Creek upstream of culvert. Q=.182(stage-6.5)^ 1.3139. The rating curve for downstream of the culvert is 0.0245*(stage -10).^1.7951.
Budd Creek